## startup
Orpheus RAS 9 SOFS-7.5 | RAS3-48-500FH | serial no: 14384-01
DEPLOYMENT: | IN2018_V07 22/02/2018 | reconfigured to fit in SOFS 7.5 buoy.
Date deployment | 22 August 2018, toffee bungs on at 1240 AET and in the water at 1310, anchor release 20:32 UTC+10.
Date first sample | 27th August 2018 see log
RECOVERY | IN2019_V02 23/03/2019
start (see RAS flushing.docx) | copious flushed with MQ, top of valve connected to bottom, with bypass tube to drain.
Tedlar bags | Bag rinsing: checked for fibres, rinsed three times. Washer connects to MQ with tee and syringe for emptying. | [6 empty] | (setB-1326-A Tedlar 2013) new set available with insurance replacement (Orpheus).
odd numbers | 250uL saturated mercuric chloride inserted into the bag with a long needle. Stored under fume cupboard.Increase concentration because of respiration? Additional 100uL after reprime because of losses in repriming. | [6 empty] | arrived and in cool room
even numbers | 250uL saturated mercuric chloride inserted into the bag with a long needle. Stored under fume cupboard.Increase concentration because of respiration? Additional 100uL after reprime because of losses in repriming.
prime (see RAS5_prime_poison_2013) | bags squeezed flat after 5mL boiled MQ so volume of prime less than 3mL + the volume of the feed tube only.
bubbles in the feed tubes can be chased out with a long needle or length of small id tube in kit with syringe.
all prime water is fresh MQ, boiled to degass.
MQ in outer acrylic tube, Ok fresh from MQ and still.
RAS deployed SOFS7, broke off and took two samples. Left loaded and reprimed prior to SOFS7.5. A lot of problems with leaking acrylic tubes.
small hardware
odds | polypropylene nuts: replaced with delrin.
all tops | bleed capillary: replaced with peek.
acrylic tubes | marine epoxy (Epiglue-International) used to reglue rebate that have peeled during deployment. (the bottoms can develop a leak but not separate). All new tubes this time.
L tie downs | new configuration.
acrylic caps, no filter assembly | annular priming grooves
general check | check capillary tubes for pressure compensation on assembly.
check pump magnet and pins. kit available
machine the bag spanner thickness down if the stem on the new set of bags is shorter.
| Check for leaking pressure comp oil from valve motor diaphragm. | Note that Artemis started leaking, 24th nov 2017, replaced before deplolyment
cap spanner made by workshop
Order parts for SOFS reconfiguration | quantitiy | part | supplier | cat no | description | pricing | drawing | status
| 50 | Colder connectors male (RAS supplied with these) | FlowTherm Australia - Sarah Knight - 0395881255 | SMF0212 | 1/8 hose barb non valved in line polypropylene coupling body (2.6mm id) | $232 total |  | arrived
| 50 | Colder connectors female (RAS supplied with these) | FlowTherm Australia - Sarah Knight - 0395881255 | SMM0212 | 1/8 hose barb non valved in line polypropylene coupling body (2.6mm id) | $321 total |  | arrived
| 50 | Value Plastics quick connect female | Chromtech Jim Jeffs | BPLF230-8012 | 1/8 200 series barb cuff connector | 12.5
| 50 | Value Plastics quick connect male | Chromtech Jim Jeffs | BPLM230-8012 | 1/8 200 series barb cuff connector | 142.50 total | Nordson site
| 100 | 1/4-28 for 1/8 ferrule short nuts delrin black | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | XP-308X | Idex Upchurch flat bottomed port flangeless, only black now | 10 pk $41 includes ferrules | http://adelab.com.au/oldpdf/Chapter_Fittings4.pdf
| 50 | 1/4-28 for 1/16 ferrule short nuts peek | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | xp-235x | Idex Upchurch flat bottomed port flangeless, not available in delrin | 10 pk $71 no ferrules but 1/16" is captive and reusable.
| not required | 1/4-28 for 1/8 ferrules short nuts delrin red | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | not available | Idex Upchurch flat bottomed port flangeless | not required
| 50 | 1/8-27NPT panel elbows 200 barb 1/8"id Value plastics | Chromtech Jim Jeffs | 18L230-6005 | natural polypropylene or J1A Tynar, Tynar ordered. | $91.50 total | http://www.nordsonmedical.com/technical/part_drawings/18L230.pdf | ordered, delivery December
| not required | 1/8' reverse ferrules flangeless flat bottom port ETFE yellow | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | XP300X | Idex Upchurch flat bottomed port flangeless | not required
| not required | 1/16' reverse ferrules flangeless flat bottom port ETFE blue | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | XP200x | Idex Upchurch flat bottomed port flangeless | not required
| 50m | 1/8"od 1/16" id tubing, bag to top of valve: PFA natural | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | 1509L | Idex | 50ft lengths $330 | https://www.idex-hs.com/pfa-tubing-natural-1-8-od-x-1-16-id-x-50ft.html
| suggested alternative | 1/8"od 1/16" id tubing, bag to top of valve: FEP natural | Adelab SA Keith Plomer - 0882347955 (alternative Fischersci) | 1521L |  | 50ft $260
| <50m | 1/4"od 1/8"id 85D ShoreA clear tubing, acrylic bottom fitting to valve bottom: ether urethane transparent | McMaster order with Craig? | 5195T63 | https://www.mcmaster.com/#standard-plastic-and-rubber-tubing/=1adydc3 | 100ft $63.00 USD | surprisingly hard
| suggested alternative | polyurethane 1/8x1/4x100' Durometer 80A, clear?????? | Chromtech Jim Jeffs | UROO-3062C |  | 100ft $233.67
| suggested alternative | 1/8"od 1/16" id tubing, bag to top of valve: PFA natural | Chromtech Jim Jeffs | JR-T-4005 |  | 25.86 per m
| 8 to 16m | ether polyurethane 4mm id black for intake/outlet | Frelin Wade Fre-thane abuse proof | 1J-025-01 | black to exclude light | 50ft @ $22.59USD | https://www.freelin-wade.com/fre-thane-85a-polyurethane-p-1271-l-en.html
| 50 | 1/4"od 1/8"id 85D ShoreA clear tubing, acrylic bottom fitting to valve bottom: ether urethane transparent | Freelin Wade (best source) | 1D-025-27 | clear | 100ft $50.50 USD | https://www.freelin-wade.com/fre-thane-85a-polyurethane-p-1120-l-en.html
| suggested alternative | neoprene 1/8x1/4x50' Durometer 80A, black | McMaster Carr | 5034K21 | black to exclude light |  | https://www.mcmaster.com/#standard-plastic-and-rubber-tubing/=1afzyfv
| 2x250mL | Microscopy grade glutardialdehyde 25% extra pure long expiry date shipped cool | Imbros | 1.04239.0250 | ex Germany 3 weeks | 2x$63.30
| 50 | Silicone o rings for caps BS141, check shoreA hardness (70?, seem too firm compared with Artemis, check with durometer)
|  | double headed cable ties or maybe stiff single strand insulated copper wire for sheath attachment |  |  |  | $6.99 per 25 | https://www.bunnings.com.au/smart-370mm-black-double-headed-cable-tie-25-pack_p4430494
| antifouling inlet and outlet bungs (toffee laden) | cupronickel 90:10 extras fabricated, drawing attached.
Luer valves and biggest possible syringe for priming.
| 8 | M8 316 SS eye nuts for lifting | PJ's |  |  | $8 each
Intake/outlet | Cu-Ni (80-20) alloy tube - ( Ken Johnson) with sugar bung, 90:10 better fouling resistance see Australwright | Used cupronickel bronze tube at the intake but unusual diatom with colonial jelly grew on it.
| sheated urethane tubes cable tied/ twist wire attached staggered, sheath= | Barfell blueline 12.5 1MPA WP black inside | Brierly Hose and Handling
| volume per m of 1/16" id teflon | 1.9856
Calculations | volume per m of 1/8" id hose in mL | 7.9173 | mL
| measured pump volume mL | 4 | mL
| smallest diameter in flow path mm | 1.39 | mm
| inlet hose length | 4.5 | m
| outlet hose length | 4.1 | m
| volume pump and inlet+outlet | 72.0888 | mL
| inlet volume bag to valve .84m | 1.6679 | mL
system volume (shortened)
| measured total volume with 9m of tube connected | 71 | mL
| 8.6m | 68.0888 | mL
the syringe will only push water forward ie pump direction when sampling
annecdotally the chain is clean below 2m
Breakage on recovery
Breakage on deployment
Other details | valve to bracket bolts | need to be at least 55mm for unmodified valve bracket. UNF hard to get in 316ss. 10-32 UNF
| top of pump bolts | 8-32 UNF

## notes
RAS9 on return to Hobart
RAS had retained its prime.
Note that the RAS can be awash
How much air in the outer acrylic tubes is too much? Outer tube volume >700mL
Eye bolts for lifting suffered crevice/pitting and broke on extraction from the float.
Extensive corrosion and rust deposited on all surfaces. Looks like stray current. Small amount of biofilm. Carbonate scraping taken by Ruth.
The barfell hose came back with a good growth of unusual chaining diatoms (Ruth sampled) and goose barnacles
One CTD cast at the site on deployment

## logs
Schedule Verification
Event 1 of 48 at 08/27/18 16:00:00
Event 2 of 48 at 08/27/18 17:00:00
Event 3 of 48 at 09/05/18 16:00:00
Event 4 of 48 at 09/05/18 17:00:00
Event 5 of 48 at 09/14/18 16:00:00
Event 6 of 48 at 09/14/18 17:00:00
Event 7 of 48 at 09/23/18 16:00:00
Event 8 of 48 at 09/23/18 17:00:00
Event 9 of 48 at 10/02/18 16:00:00
Event 10 of 48 at 10/02/18 17:00:00
Event 11 of 48 at 10/11/18 16:00:00
Event 12 of 48 at 10/11/18 17:00:00
Event 13 of 48 at 10/20/18 16:00:00
Event 14 of 48 at 10/20/18 17:00:00
Event 15 of 48 at 10/29/18 16:00:00
Event 16 of 48 at 10/29/18 17:00:00
Event 17 of 48 at 11/07/18 16:00:00
Event 18 of 48 at 11/07/18 17:00:00
Event 19 of 48 at 11/16/18 16:00:00
Event 20 of 48 at 11/16/18 17:00:00
Event 21 of 48 at 11/25/18 16:00:00
Event 22 of 48 at 11/25/18 17:00:00
Event 23 of 48 at 12/04/18 16:00:00
Event 24 of 48 at 12/04/18 17:00:00
Event 25 of 48 at 12/13/18 16:00:00
Event 26 of 48 at 12/13/18 17:00:00
Event 27 of 48 at 12/22/18 16:00:00
Event 28 of 48 at 12/22/18 17:00:00
Event 29 of 48 at 12/31/18 16:00:00
Event 30 of 48 at 12/31/18 17:00:00
Event 31 of 48 at 01/09/19 16:00:00
Event 32 of 48 at 01/09/19 17:00:00
Event 33 of 48 at 01/18/19 16:00:00
Event 34 of 48 at 01/18/19 17:00:00
Event 35 of 48 at 01/27/19 16:00:00
Event 36 of 48 at 01/27/19 17:00:00
Event 37 of 48 at 02/05/19 16:00:00
Event 38 of 48 at 02/05/19 17:00:00
Event 39 of 48 at 02/14/19 16:00:00
Event 40 of 48 at 02/14/19 17:00:00
Event 41 of 48 at 02/23/19 16:00:00
Event 42 of 48 at 02/23/19 17:00:00
Event 43 of 48 at 03/04/19 16:00:00
Event 44 of 48 at 03/04/19 17:00:00
Event 45 of 48 at 03/13/19 16:00:00
Event 46 of 48 at 03/13/19 17:00:00
Event 47 of 48 at 03/22/19 16:00:00
Event 48 of 48 at 03/22/19 17:00:00
Modify an event [N] ?
Header A¦ SOFS-7.5 RAS in Float
B¦ Prep Di Davies 2018-07-25
C¦
Acid D¦ Pre-sample acid flush: Disabled
E¦ Flushing volume = NA [ml]
F¦ Flushing time limit = NA [min]
G¦ Exposure time delay = NA [min]
Water H¦ Flushing volume = 100 [ml]
I¦ Flushing time limit = 5 [min]
Sample J¦ Sample volume = 500 [ml]
K¦ Sample time limit = 25 [min]
Acid L¦ Post-sample acid flush: Enabled
M¦ Flushing volume = 10 [ml]
N¦ Flushing time limit = 1 [min]
Timing P¦ Pump data period = 15 [sec]
V¦ Verify and proceed.
Selection [‰] ? c
> Take #2, deployment 2018-08-22
Header A¦ SOFS-7.5 RAS in Float
B¦ Prep Di Davies 2018-07-25
C¦ Take #2, deployment 2018-08-22
Acid D¦ Pre-sample acid flush: Disabled
E¦ Flushing volume = NA [ml]
F¦ Flushing time limit = NA [min]
G¦ Exposure time delay = NA [min]
Water H¦ Flushing volume = 100 [ml]
I¦ Flushing time limit = 5 [min]
Sample J¦ Sample volume = 500 [ml]
K¦ Sample time limit = 25 [min]
Acid L¦ Post-sample acid flush: Enabled
M¦ Flushing volume = 10 [ml]
N¦ Flushing time limit = 1 [min]
Timing P¦ Pump data period = 15 [sec]
V¦ Verify and proceed.
Selection [C] ? v
Performing 6 second low-power sleep mode test...
System status:
Date Time Battery Temp Port
08/17/18 01:19:52 37.5 Vb 12.0∞C 00 (home)
Caution: Deployment will overwrite the
EEPROM data backup cache.
Proceed with the deployment [N] ? y
Erasing EEPROM entries .........................
>>> Remove communication cable and attach dummy plug. <<<
>>> System is ready to deploy. <<<
08/17/18 01:20:06 Waiting for event 01 of 48 @ 08/27/18 16:00:00
08/17/18 01:20:07 Suspended until 08/27/18 16:00:00 ...

## bag wts
data file: IMOS_ABOS-SOTS_FRVMWEBUOSTCGP_20180628_SOFS_FV02_SOFS-7.5-2018-Gridded-Data_END-20190423_C-20190916.nc
RAS and mooring | Sample | total wt g | net sample wt g. |  |  |  | date time utc mm/dd/yy hh:mm:ss | 25075.6667 | 9.8006 | 34.7138 | 1026.769 | 2018-08-27 16:00 | 16 | sorted data
RAS 9 SOFS7.5 # | Event 1 of 48 at 08/27/18 16:00:00 | 428.86 | 418.553 |  |  |  | 08/27/18 16:00:00 | 25084.6667 | 9.4345 | 34.6357 | 1026.7684 | 2018-09-05 16:00 | 16
RAS 9 SOFS7.5 # | Event 2 of 48 at 08/27/18 17:00:00 | 430.15 | 419.843 |  |  |  | 08/27/18 17:00:00 | 25093.6667 | 9.4449 | 34.6216 | 1026.756 | 2018-09-14 16:00 | 16
RAS 9 SOFS7.5 # | Event 3 of 48 at 09/05/18 16:00:00 | 411.38 | 401.073 |  |  |  | 09/05/18 16:00:00 | 25102.6667 | 9.6593 | 34.6872 | 1026.7717 | 2018-09-23 16:00 | 16
RAS 9 SOFS7.5 # | Event 4 of 48 at 09/05/18 17:00:00 | 401.26 | 390.953 |  |  |  | 09/05/18 17:00:00 | 25111.6667 | 9.3318 | 34.6211 | 1026.7739 | 2018-10-02 16:00 | 16
RAS 9 SOFS7.5 # | Event 5 of 48 at 09/14/18 16:00:00 | 430.54 | 420.233 |  |  |  | 09/14/18 16:00:00 | 25120.6667 | 8.998 | 34.5037 | 1026.7362 | 2018-10-11 16:00 | 16
RAS 9 SOFS7.5 # | Event 6 of 48 at 09/14/18 17:00:00 | 417.54 | 407.233 |  |  |  | 09/14/18 17:00:00 | 25129.6667 | 9.6384 | 34.5548 | 1026.6718 | 2018-10-20 16:00 | 16
RAS 9 SOFS7.5 # | Event 7 of 48 at 09/23/18 16:00:00 | 415.65 | 405.343 |  |  |  | 09/23/18 16:00:00 | 25138.6667 | 9.1461 | 34.5446 | 1026.7443 | 2018-10-29 16:00 | 16
RAS 9 SOFS7.5 # | Event 8 of 48 at 09/23/18 17:00:00 | 413.02 | 402.713 |  |  |  | 09/23/18 17:00:00 | 25147.6667 | 9.1443 | 34.5593 | 1026.7566 | 2018-11-07 16:00 | 16
RAS 9 SOFS7.5 # | Event 9 of 48 at 10/02/18 16:00:00 | 423.34 | 413.033 |  |  |  | 10/02/18 16:00:00 | 25156.6667 | 9.4704 | 34.5695 | 1026.711 | 2018-11-16 16:00 | 16
RAS 9 SOFS7.5 # | Event 10 of 48 at 10/02/18 17:00:00 | 425.34 | 415.033 |  |  |  | 10/02/18 17:00:00 | 25165.6667 | 9.9274 | 34.5519 | 1026.6206 | 2018-11-25 16:00 | 16
RAS 9 SOFS7.5 # | Event 11 of 48 at 10/11/18 16:00:00 | 417.09 | 406.783 |  |  |  | 10/11/18 16:00:00 | 25174.6667 | 9.8432 | 34.5418 | 1026.6274 | 2018-12-04 16:00 | 16
RAS 9 SOFS7.5 # | Event 12 of 48 at 10/11/18 17:00:00 | 415.48 | 405.173 |  |  |  | 10/11/18 17:00:00 | 25183.6667 | 10.4065 | 34.5264 | 1026.5189 | 2018-12-13 16:00 | 16
RAS 9 SOFS7.5 # | Event 13 of 48 at 10/20/18 16:00:00 | 424.97 | 414.663 |  |  |  | 10/20/18 16:00:00 | 25192.6667 | 11.5403 | 34.5784 | 1026.3542 | 2018-12-22 16:00 | 16
RAS 9 SOFS7.5 # | Event 14 of 48 at 10/20/18 17:00:00 | 363.3 | 352.993 |  |  |  | 10/20/18 17:00:00 | 25201.6667 | 10.7488 | 34.5436 | 1026.4716 | 2018-12-31 16:00 | 16
RAS 9 SOFS7.5 # | Event 15 of 48 at 10/29/18 16:00:00 | 402.34 | 392.033 |  |  |  | 10/29/18 16:00:00 | 25210.6667 | 10.996 | 34.5599 | 1026.4401 | 2019-01-09 16:00 | 16
RAS 9 SOFS7.5 # | Event 16 of 48 at 10/29/18 17:00:00 | 409.36 | 399.053 |  |  |  | 10/29/18 17:00:00 | 25219.6667 | 11.758 | 34.6129 | 1026.341 | 2019-01-18 16:00 | 16
RAS 9 SOFS7.5 # | Event 17 of 48 at 11/07/18 16:00:00 | 370.48 | 360.173 |  |  |  | 11/07/18 16:00:00 | 25228.6667 | 10.9153 | 34.5425 | 1026.4407 | 2019-01-27 16:00 | 16
RAS 9 SOFS7.5 # | Event 18 of 48 at 11/07/18 17:00:00 | 410.93 | 400.623 |  |  |  | 11/07/18 17:00:00 | 25237.6667 | 11.3754 | 34.5928 | 1026.3962 | 2019-02-05 16:00 | 16
RAS 9 SOFS7.5 # | Event 19 of 48 at 11/16/18 16:00:00 | 423.9 | 413.593 |  |  |  | 11/16/18 16:00:00 | 25246.6667 | 10.5504 | 34.5159 | 1026.4847 | 2019-02-14 16:00 | 16
RAS 9 SOFS7.5 # | Event 20 of 48 at 11/16/18 17:00:00 | 378.66 | 368.353 |  |  |  | 11/16/18 17:00:00 | 25255.6667 | 10.82 | 34.534 | 1026.4514 | 2019-02-23 16:00 | 16
RAS 9 SOFS7.5 # | Event 21 of 48 at 11/25/18 16:00:00 | 379.69 | 369.383 |  |  |  | 11/25/18 16:00:00 | 25264.6667 | 11.4233 | 34.5692 | 1026.3693 | 2019-03-04 16:00 | 16
RAS 9 SOFS7.5 # | Event 22 of 48 at 11/25/18 17:00:00 | 418.18 | 407.873 |  |  |  | 11/25/18 17:00:00 | 25273.6667 | 10.5746 | 34.5506 | 1026.5079 | 2019-03-13 16:00 | 16
RAS 9 SOFS7.5 # | Event 23 of 48 at 12/04/18 16:00:00 | 423.9 | 413.593 |  |  |  | 12/04/18 16:00:00 | 25282.6667 | 11.5177 | 34.5655 | 1026.3488 | 2019-03-22 16:00 | 16
RAS 9 SOFS7.5 # | Event 24 of 48 at 12/04/18 17:00:00 | 425.24 | 414.933 |  |  |  | 12/04/18 17:00:00 | 25075.7083 | 9.793 | 34.7147 | 1026.7693 | 2018-08-27 17:00 | 17
RAS 9 SOFS7.5 # | Event 25 of 48 at 12/13/18 16:00:00 | 350.13 | 339.823 |  |  |  | 12/13/18 16:00:00 | 25084.7083 | 9.4969 | 34.648 | 1026.7683 | 2018-09-05 17:00 | 17
RAS 9 SOFS7.5 # | Event 26 of 48 at 12/13/18 17:00:00 | 394.48 | 384.173 |  |  |  | 12/13/18 17:00:00 | 25093.7083 | 9.4426 | 34.6227 | 1026.757 | 2018-09-14 17:00 | 17
RAS 9 SOFS7.5 # | Event 27 of 48 at 12/22/18 16:00:00 | 406.61 | 396.303 |  |  |  | 12/22/18 16:00:00 | 25102.7083 | 9.6511 | 34.6862 | 1026.7725 | 2018-09-23 17:00 | 17
RAS 9 SOFS7.5 # | Event 28 of 48 at 12/22/18 17:00:00 | 420.63 | 410.323 |  |  |  | 12/22/18 17:00:00 | 25111.7083 | 9.3273 | 34.6216 | 1026.7758 | 2018-10-02 17:00 | 17
RAS 9 SOFS7.5 # | Event 29 of 48 at 12/31/18 16:00:00 | 407.08 | 396.773 |  |  |  | 12/31/18 16:00:00 | 25120.7083 | 8.9429 | 34.4945 | 1026.7377 | 2018-10-11 17:00 | 17
RAS 9 SOFS7.5 # | Event 30 of 48 at 12/31/18 17:00:00 | 422.93 | 412.623 |  |  |  | 12/31/18 17:00:00 | 25129.7083 | 9.6268 | 34.5552 | 1026.6736 | 2018-10-20 17:00 | 17
RAS 9 SOFS7.5 # | Event 31 of 48 at 01/09/19 16:00:00 | 414.04 | 403.733 |  |  |  | 01/09/19 16:00:00 | 25138.7083 | 9.1446 | 34.5448 | 1026.7448 | 2018-10-29 17:00 | 17
RAS 9 SOFS7.5 # | Event 32 of 48 at 01/09/19 17:00:00 | 428.93 | 418.623 |  |  |  | 01/09/19 17:00:00 | 25147.7083 | 9.1436 | 34.56 | 1026.757 | 2018-11-07 17:00 | 17
RAS 9 SOFS7.5 # | Event 33 of 48 at 01/18/19 16:00:00 | 421.9 | 411.593 |  |  |  | 01/18/19 16:00:00 | 25156.7083 | 9.4662 | 34.5688 | 1026.7108 | 2018-11-16 17:00 | 17
RAS 9 SOFS7.5 # | Event 34 of 48 at 01/18/19 17:00:00 | 413.48 | 403.173 |  |  |  | 01/18/19 17:00:00 | 25165.7083 | 9.9175 | 34.5505 | 1026.6213 | 2018-11-25 17:00 | 17
RAS 9 SOFS7.5 # | Event 35 of 48 at 01/27/19 16:00:00 | 426.21 | 415.903 |  |  |  | 01/27/19 16:00:00 | 25174.7083 | 9.848 | 34.5413 | 1026.626 | 2018-12-04 17:00 | 17
RAS 9 SOFS7.5 # | Event 36 of 48 at 01/27/19 17:00:00 | 432.21 | 421.903 |  |  |  | 01/27/19 17:00:00 | 25183.7083 | 10.4151 | 34.5264 | 1026.5173 | 2018-12-13 17:00 | 17
RAS 9 SOFS7.5 # | Event 37 of 48 at 02/05/19 16:00:00 | 421.86 | 411.553 |  |  |  | 02/05/19 16:00:00 | 25192.7083 | 11.5288 | 34.572 | 1026.3546 | 2018-12-22 17:00 | 17
RAS 9 SOFS7.5 # | Event 38 of 48 at 02/05/19 17:00:00 | 422.46 | 412.153 |  |  |  | 02/05/19 17:00:00 | 25201.7083 | 10.7401 | 34.5446 | 1026.4738 | 2018-12-31 17:00 | 17
RAS 9 SOFS7.5 # | Event 39 of 48 at 02/14/19 16:00:00 | 414.13 | 403.823 |  |  |  | 02/14/19 16:00:00 | 25210.7083 | 10.9866 | 34.5592 | 1026.4409 | 2019-01-09 17:00 | 17
RAS 9 SOFS7.5 # | Event 40 of 48 at 02/14/19 17:00:00 | 337.55 | 327.243 |  |  |  | 02/14/19 17:00:00 | 25219.7083 | 11.762 | 34.611 | 1026.3383 | 2019-01-18 17:00 | 17
RAS 9 SOFS7.5 # | Event 41 of 48 at 02/23/19 16:00:00 | 417.93 | 407.623 |  |  |  | 02/23/19 16:00:00 | 25228.7083 | 10.8936 | 34.5397 | 1026.4423 | 2019-01-27 17:00 | 17
RAS 9 SOFS7.5 # | Event 42 of 48 at 02/23/19 17:00:00 | 424.57 | 414.263 |  |  |  | 02/23/19 17:00:00 | 25237.7083 | 11.3633 | 34.5929 | 1026.3983 | 2019-02-05 17:00 | 17
RAS 9 SOFS7.5 # | Event 43 of 48 at 03/04/19 16:00:00 | 396.02 | 385.713 |  |  |  | 03/04/19 16:00:00 | 25246.7083 | 10.5432 | 34.516 | 1026.4862 | 2019-02-14 17:00 | 17
RAS 9 SOFS7.5 # | Event 44 of 48 at 03/04/19 17:00:00 | 383.35 | 373.043 |  |  |  | 03/04/19 17:00:00 | 25255.7083 | 10.8014 | 34.5286 | 1026.4504 | 2019-02-23 17:00 | 17
RAS 9 SOFS7.5 # | Event 45 of 48 at 03/13/19 16:00:00 | 408.25 | 397.943 |  |  |  | 03/13/19 16:00:00 | 25264.7083 | 11.4328 | 34.5696 | 1026.3676 | 2019-03-04 17:00 | 17
RAS 9 SOFS7.5 # | Event 46 of 48 at 03/13/19 17:00:00 | 419.43 | 409.123 |  |  |  | 03/13/19 17:00:00 | 25273.7083 | 10.5682 | 34.5513 | 1026.5092 | 2019-03-13 17:00 | 17
RAS 9 SOFS7.5 # | Event 47 of 48 at 03/22/19 16:00:00 | 399.77 | 389.463 |  |  |  | 03/22/19 16:00:00 | 25282.7083 | 11.5099 | 34.5647 | 1026.3496 | 2019-03-22 17:00 | 17
RAS 9 SOFS7.5 # | Event 48 of 48 at 03/22/19 17:00:00 | 419 | net sample wt g. | sample analysis |  |  | date time utc mm/dd/yy hh:mm:ss | time since 1950 01 01 | temp C 1.51m SBE37SM 7408 | PSAL 1.51m SBE 37SM 7408 | density kg m3 Asimet L22 | date time | hour (sort)
Bag batch 2014 - 5
empty with cap.
10.24
10.45
10.25
10.3
10.36
10.32
10.33
10.32
10.18
10.32
| mean | 10.307
cast 1 | end posn lat dd mm.mm | long dd.mm.mm | date time | psal 1 | psal 2 | temp 1 | temp 2 | pressure | density
| 47 01.279 | 142 11.918 | 2018-08-22 13:09 | 34.7682 | 34.7709 | 10.0257 | 10.0262 | 5.154 | 1026.7666
|  |  | 2018-08-22 13:09 | 34.7682 | 34.7709 | 10.026 | 10.0265 | 5.7539 | 1026.7665
|  |  | 2018-08-22 13:09 | 34.7682 | 34.7708 | 10.0258 | 10.0264 | 5.9235 | 1026.7665
deployment cast file:
Note on C50: Microsoft Office User: Microsoft Office User: estimate
Note on K50: Microsoft Office User: comment in nc file that salinity bad after 2019-02-14, comparision with bottle data, no specific senor mnetioned.

## diln
|  | SOMMA | data file: IMOS_ABOS-SOTS_FRVMWEBUOSTCGP_20180628_SOFS_FV02_SOFS-7.5-2018-Gridded-Data_END-20190423_C-20190916.nc
| #RAS | salinity | temp C 1.51m SBE37SM 7408 | PSAL 1.51m SBE 37SM 7408 | density kg m3 Asimet L22 | diln | net sample g | approximate volume of prime | net sample kg
2018-08-27 17:00 | 2 | 34.142 | 9.793 | 34.7147 | 1026.7693 | 0.9835 | 414.933 | 6.8453 | 0.41493
2018-09-05 17:00 | 4 | 34.264 | 9.4969 | 34.648 | 1026.7683 | 0.98892 | 339.823 | 3.7662 | 0.33982
2018-09-14 17:00 | 6 | 34.323 | 9.4426 | 34.6227 | 1026.757 | 0.99134 | 384.173 | 3.3255 | 0.38417
2018-09-23 17:00 | 8 | 34.323 | 9.6511 | 34.6862 | 1026.7725 | 0.98953 | 396.303 | 4.1497 | 0.3963
2018-10-02 17:00 | 10 | 34.215 | 9.3273 | 34.6216 | 1026.7758 | 0.98826 | 410.323 | 4.8189 | 0.41032
2018-10-11 17:00 | 12 | 33.738 | 8.9429 | 34.4945 | 1026.7377 | 0.97807 | 396.773 | 8.7016 | 0.39677
2018-10-20 17:00 | 14 | 34.202 | 9.6268 | 34.5552 | 1026.6736 | 0.98978 | 412.623 | 4.2176 | 0.41262
2018-10-29 17:00 | 16 | 34.146 | 9.1446 | 34.5448 | 1026.7448 | 0.98846 | 403.733 | 4.6609 | 0.40373
2018-11-07 17:00 | 18 | 34.191 | 9.1436 | 34.56 | 1026.757 | 0.98932 | 418.623 | 4.4697 | 0.41862
2018-11-16 17:00 | 20 | 34.197 | 9.4662 | 34.5688 | 1026.7108 | 0.98924 | 411.593 | 4.4268 | 0.41159
2018-11-25 17:00 | 22 | 34.171 | 9.9175 | 34.5505 | 1026.6213 | 0.98902 | 403.173 | 4.4284 | 0.40317
2018-12-04 17:00 | 24 | 34.105 | 9.848 | 34.5413 | 1026.626 | 0.98737 | 415.903 | 5.2534 | 0.4159
2018-12-13 17:00 | 26 | 34.026 | 10.4151 | 34.5264 | 1026.5173 | 0.98551 | 421.903 | 6.1147 | 0.4219
2018-12-22 17:00 | 28 | 33.742 | 11.5288 | 34.572 | 1026.3546 | 0.97599 | 411.553 | 9.8805 | 0.41155
2018-12-31 17:00 | 30 | 33.544 | 10.7401 | 34.5446 | 1026.4738 | 0.97103 | 412.153 | 11.9382 | 0.41215
2019-01-09 17:00 | 32 | 34.191 | 10.9866 | 34.5592 | 1026.4409 | 0.98935 | 403.823 | 4.3024 | 0.40382
2019-01-18 17:00 | 34 | 33.908 | 11.762 | 34.611 | 1026.3383 | 0.97969 | 327.243 | 6.6468 | 0.32724
2019-01-27 17:00 | 36 | 34.204 | 10.8936 | 34.5397 | 1026.4423 | 0.99028 | 407.623 | 3.9618 | 0.40762
2019-02-05 17:00 | 38 | 34.225 | 11.3633 | 34.5929 | 1026.3983 | 0.98936 | 414.263 | 4.4057 | 0.41426
2019-02-14 17:00 | 40 | 34.122 | 10.5432 | 34.516 | 1026.4862 | 0.98859 | 385.713 | 4.4029 | 0.38571
2019-02-23 17:00 | 42 | 33.892 | 10.8014 | 34.5286 | 1026.4504 | 0.98156 | 373.043 | 6.8778 | 0.37304
2019-03-04 17:00 | 44 | 34.217 | 11.4328 | 34.5696 | 1026.3676 | 0.9898 | 397.943 | 4.0589 | 0.39794
2019-03-13 17:00 | 46 | 34.179 | 10.5682 | 34.5513 | 1026.5092 | 0.98922 | 409.123 | 4.4084 | 0.40912
2019-03-22 17:00 | 48 | 34.217 | 11.5099 | 34.5647 | 1026.3496 | 0.98994 | 389.463 | 3.9178 | 0.38946
2018-08-22 13:09 | ctd IN2019_07 |  | 10.0257 | 34.7682 | 1026.7666
2018-08-22 13:09 | ctd IN2019_07 |  | 10.026 | 34.7682 | 1026.7665
2018-08-22 13:09 | ctd IN2019_07 |  | 10.0258 | 34.7682 | 1026.7665
Note on E2: Microsoft Office User: comment in nc file that salinity bad after 2019-02-14, comparision with bottle data, no specific senor mnetioned.

## plots
CTD | date time | ctd psal 1 | ctd temp 2
| 2018-08-22 13:09 | 34.7682 | 10.0262
| 2018-08-22 13:09 | 34.7682 | 10.0265
| 2018-08-22 13:09 | 34.7682 | 10.0264
| date/time | NOx | Phosphate | Silicate
| ddmmyy hh:mm | umol Kg-1 | umol Kg-1 | umol Kg-1 |  | N/P | N/Si
RAS 2 | 2018-08-27 17:00 | 12.0169 | 0.95561 | 3.0698 | 2 | 12.5751 | 3.9145
RAS 4 | 2018-09-05 17:00 | 12.6257 | 1.0242 | 2.5606 | 4 | 12.3269 | 4.9308
RAS 6 | 2018-09-14 17:00 | 12.5163 | 1.0905 | 2.5544 | 6 | 11.4775 | 4.9
RAS 8 | 2018-09-23 17:00 | 12.2734 | 0.98423 | 2.3622 | 8 | 12.47 | 5.1958
RAS 10 | 2018-10-02 17:00 | 12.7819 | 1.0052 | 2.2666 | 10 | 12.7157 | 5.6391
RAS 12 | 2018-10-11 17:00 | 13.3437 | 1.0655 | 3.2861 | 12 | 12.5234 | 4.0606
RAS 14 | 2018-10-20 17:00 | 12.5765 | 1.2104 | 3.7395 | 14 | 10.3902 | 3.3632
RAS 16 | 2018-10-29 17:00 | 13.6074 | 1.212 | 3.7442 | 16 | 11.2276 | 3.6342
RAS 18 | 2018-11-07 17:00 | 11.6461 | 0.93523 | 3.0518 | 18 | 12.4526 | 3.8161
RAS 20 | 2018-11-16 17:00 | 10.7811 | 0.94519 | 2.7568 | 20 | 11.4062 | 3.9107
RAS 22 | 2018-11-25 17:00 | 11.3853 | 1.0735 | 2.7577 | 22 | 10.6055 | 4.1286
RAS 24 | 2018-12-04 17:00 | 12.1737 | 1.3121 | 5.0313 | 24 | 9.2782 | 2.4196
RAS 26 | 2018-12-13 17:00 | 10.913 | 1.1121 | 3.8057 | 26 | 9.8133 | 2.8675
RAS 28 | 2018-12-22 17:00 | 7.6369 | 0.82858 | 1.2978 | 28 | 9.2169 | 5.8846
RAS 30 | 2018-12-31 17:00 | 10.745 | 1.0635 | 2.5082 | 30 | 10.1038 | 4.284
RAS 32 | 2019-01-09 17:00 | 9.7045 | 0.90595 | 1.5756 | 32 | 10.712 | 6.1594
RAS 34 | 2019-01-18 17:00 | 9.1498 | 0.97465 | 1.2929 | 34 | 9.3878 | 7.0769
RAS 36 | 2019-01-27 17:00 | 9.7298 | 0.88542 | 1.1806 | 36 | 10.9889 | 8.2417
RAS 38 | 2019-02-05 17:00 | 9.5915 | 0.92567 | 1.6741 | 38 | 10.3617 | 5.7294
RAS 40 | 2019-02-14 17:00 | 10.8498 | 1.1333 | 2.6607 | 40 | 9.5739 | 4.0778
RAS 42 | 2019-02-23 17:00 | 9.0618 | 0.8635 | 1.2903 | 42 | 10.4943 | 7.0231
RAS 44 | 2019-03-04 17:00 | 7.9339 | 0.78748 | 0.88591 | 44 | 10.075 | 8.9556
RAS 46 | 2019-03-13 17:00 | 10.2516 | 1.1916 | 3.1513 | 46 | 8.6033 | 3.2531
RAS 48 | 2019-03-22 17:00 | 8.7645 | 1.0777 | 2.0177 | 48 | 8.1324 | 4.3439
CTD | 2018-10-01 11:59 | 12.3494 | 0.93497 | 2.4348 |  | 13.2083 | 5.072
CTD | 2018-10-01 11:59 | 12.398 | 0.93497 | 2.4348 |  | 13.2604 | 5.092
CTD | 2018-08-22 11:08 | 11.2957 | 0.86193 | 2.3783 |  | 13.1051 | 4.7494
CTD | 2019-03-16 15:50 | 8.7321 | 0.71315 | 0.82323 |  | 12.2445 | 10.6071
CTD | 2019-03-16 15:50 | 8.7263 | 0.72678 | 0.81154 |  | 12.0067 | 10.7527
CTD | 2019-03-17 01:45 | 9.0243 | 0.73359 | 0.93526 |  | 12.3015 | 9.649
CTD | 2019-03-19 01:35 | 9.0855 | 0.72287 | 0.99467 |  | 12.5687 | 9.1342
CTD | 2019-03-19 12:39 | 9.0343 | 0.70926 | 0.84956 |  | 12.7376 | 10.6342
CTD | 2019-03-20 06:00 | 8.5662 | 0.70443 | 0.7551 |  | 12.1604 | 11.3445
CTD | 2019-03-20 06:00 | 8.574 | 0.70248 | 0.82232 | recovery 23/03/2019 | 12.2053 | 10.4265
CTD | 2019-03-24 02:01 | 8.2395 | 0.68297 | 0.55631 |  | 12.0642 | 14.8109
CTD | 2019-03-27 01:56 | 10.9579 | 0.84456 | 1.9921 |  | 12.9746 | 5.5007
CTD | 2019-03-28 12:41 | 12.9958 | 0.96244 | 2.9448 |  | 13.503 | 4.4132
CTD | 2019-03-29 01:34 | 10.5733 | 0.82315 | 1.6132 |  | 12.845 | 6.5543
CTD | 2019-03-30 02:07 | 10.542 | 0.82217 | 1.468 |  | 12.8223 | 7.1812
CTD | 2019-03-30 13:10 | 10.8868 | 0.83482 | 1.7203 |  | 13.0408 | 6.3284
| date/time | Alkalinity | TCO2
| ddmmyy hh:mm | µmol/kg | µmol/kg
RAS 2 | 2018-08-27 17:00 | 2295.64 | 2387.7871
RAS 4 | 2018-09-05 17:00 | 2292.5607 | 2443.6173
RAS 6 | 2018-09-14 17:00 | 2289.4939 | 2344.1463
RAS 8 | 2018-09-23 17:00 | 2289.0127 | 2429.9187
RAS 10 | 2018-10-02 17:00 | 2290.1146 | 2393.6867
RAS 12 | 2018-10-11 17:00 | 2226.772 | 2294.9762
RAS 14 | 2018-10-20 17:00 | 2288.9259 | 2145.909
RAS 16 | 2018-10-29 17:00 | 2284.9939 | 2374.6793
RAS 18 | 2018-11-07 17:00 | 2239.651 | 2265.1502
RAS 20 | 2018-11-16 17:00 | 2241.648 | 2349.2419
RAS 22 | 2018-11-25 17:00 | 2226.8459 | 2363.8089
RAS 24 | 2018-12-04 17:00 | 2283.5106 | 2216.9883
RAS 26 | 2018-12-13 17:00 | 2238.957 | 2310.0552
RAS 28 | 2018-12-22 17:00 | 2229.6512 | 2248.4044
RAS 30 | 2018-12-31 17:00 | 2247.2472 | 2295.4436
RAS 32 | 2019-01-09 17:00 | 2276.3523 | 2340.9155
RAS 34 | 2019-01-18 17:00 | 2287.9736 | 2208.8091
RAS 36 | 2019-01-27 17:00 | 2247.4053 | 2174.2773
RAS 38 | 2019-02-05 17:00 | 2293.8819 | 2238.4411
RAS 40 | 2019-02-14 17:00 | 2284.4799 | 2259.5781
RAS 42 | 2019-02-23 17:00 | 2234.3563 | 2217.5192
RAS 44 | 2019-03-04 17:00 | 2255.2434 | 2120.0691
RAS 46 | 2019-03-13 17:00 | 2287.7805 | 2112.1035
RAS 48 | 2019-03-22 17:00 | 2287.764 | 2136.6181 | depth | nitrate
IN2018_V07 | 2018-08-22 10:52 | 2301.6326 | 2115.0464 | 29.4 | 11.2957
IN2018_V07 | 2018-08-22 10:52 | 2302.1243 | 2114.8363 | 5.9 | 11.2957
IN2018_V04 | 2018-10-01 07:51 | 2296.0066 | 2117.7266 | 30 | 12.3494
IN2018_V04 | 2018-10-01 07:51 | 2295.925 | 2114.7054 | 5 | 12.398
IN2019_V02 | 2019-03-19 12:00 | 2294.4468 | 2092.4412 | 7.6 | 9.0855
IN2019_V02 | 2019-03-19 12:00 | 2295.893 | 2091.2207 | 20.7 | 9.0343
IN2019_V02 | 2019-03-24 01:22 | 2295.1525 | 2085.618 | 6.3 | 8.2395
IN2019_V02 | 2019-03-24 01:22 | 2295.9698 | 2088.1742 | 20.1 | 8.2395
Silicate
umol Kg-1 | N/Si
0.88591 | 8.9556
1.1806 | 8.2417
1.2903 | 7.0231
1.2929 | 7.0769
1.2978 | 5.8846
1.5756 | 6.1594
1.6741 | 5.7294
2.0177 | 4.3439
2.2666 | 5.6391
2.3622 | 5.1958
2.5082 | 4.284
2.5544 | 4.9
2.5606 | 4.9308
2.6607 | 4.0778
2.7568 | 3.9107
2.7577 | 4.1286
3.0518 | 3.8161
3.0698 | 3.9145
3.1513 | 3.2531
3.2861 | 4.0606
3.7395 | 3.3632
3.7442 | 3.6342
3.8057 | 2.8675
5.0313 | 2.4196

## nutrients
SOTS site underway supply (CO2)
Samples of GF/F filtered underway seawater for sediment trap processing to hydrochem, 22/08/2018. Filter was pea greeen and smelt of cod liver.
Collected underway sea water 2x10 L: |  |  | -47.02 | 142.25
Sample ID | NOx (uM) | Phosphate (uM) | Silicate (uM) | Ammonia (uM) | Nitrite (uM) | TSG salinity
DD 1 | 12.13 | 0.92 | 2.5 | 0.06 | 0.095 | 34.726
DD 2 | 12.07 | 0.91 | 2.6 | 0.06 | 0.095 | 34.74 | [6 empty] | mean results, diln correction an units conversion
CSIRO Hydrochemistry Data Report - 95 | [11 empty] | date/time | RAS# | NOx (µM) | NOx | Phosphate (µM) | Phosphate | Silicate (µM) | Silicate | diln applied | kg/m3
Lab No. | Site ID | Nutrient ID | NOx (µM) | Phosphate (µM) | Silicate (µM) | RAS# | [5 empty] | ddmmyy hh:mm |  | mean in pink | umol Kg-1 | mean in pink | umol Kg-1 | mean in pink | umol Kg-1
16743 | SOTS 7.5 | 1 | 12.11 | 0.95 | 3.2 | 2 | [5 empty] | 2018-08-27 17:00 | 2 | 12.135 | 12.0169 | 0.965 | 0.95561 | 3.1 | 3.0698 | 0.9835 | 1026.7693
16770 | SOTS 7.5 | 28 | 12.16 | 0.98 | 3 | 2 | 0.41118 |  |  |  |  | 2018-09-05 17:00 | 4 | 12.82 | 12.6257 | 1.04 | 1.0242 | 2.6 | 2.5606 | 0.98892 | 1026.7683
16744 | SOTS 7.5 | 2 | 12.82 | 1.04 | 2.6 | 4 | [5 empty] | 2018-09-14 17:00 | 6 | 12.74 | 12.5163 | 1.11 | 1.0905 | 2.6 | 2.5544 | 0.99134 | 1026.757
16745 | SOTS 7.5 | 3 | 12.74 | 1.11 | 2.6 | 6 | [5 empty] | 2018-09-23 17:00 | 8 | 12.47 | 12.2734 | 1 | 0.98423 | 2.4 | 2.3622 | 0.98953 | 1026.7725
16746 | SOTS 7.5 | 4 | 12.47 | 1 | 2.4 | 8 | [5 empty] | 2018-10-02 17:00 | 10 | 12.97 | 12.7819 | 1.02 | 1.0052 | 2.3 | 2.2666 | 0.98826 | 1026.7758
16747 | SOTS 7.5 | 5 | 12.97 | 1.02 | 2.3 | 10 | [5 empty] | 2018-10-11 17:00 | 12 | 13.4 | 13.3437 | 1.07 | 1.0655 | 3.3 | 3.2861 | 0.97807 | 1026.7377
16748 | SOTS 7.5 | 6 | 13.4 | 1.07 | 3.3 | 12 | [5 empty] | 2018-10-20 17:00 | 14 | 12.78 | 12.5765 | 1.23 | 1.2104 | 3.8 | 3.7395 | 0.98978 | 1026.6736
16749 | SOTS 7.5 | 7 | 12.78 | 1.23 | 3.8 | 14 | [5 empty] | 2018-10-29 17:00 | 16 | 13.81 | 13.6074 | 1.23 | 1.212 | 3.8 | 3.7442 | 0.98846 | 1026.7448
16750 | SOTS 7.5 | 8 | 13.81 | 1.23 | 3.8 | 16 | [5 empty] | 2018-11-07 17:00 | 18 | 11.83 | 11.6461 | 0.95 | 0.93523 | 3.1 | 3.0518 | 0.98932 | 1026.757
16751 | SOTS 7.5 | 9 | 11.83 | 0.95 | 3.1 | 18 | [5 empty] | 2018-11-16 17:00 | 20 | 10.95 | 10.7811 | 0.96 | 0.94519 | 2.8 | 2.7568 | 0.98924 | 1026.7108
16752 | SOTS 7.5 | 10 | 10.95 | 0.96 | 2.8 | 20 | [5 empty] | 2018-11-25 17:00 | 22 | 11.56 | 11.3853 | 1.09 | 1.0735 | 2.8 | 2.7577 | 0.98902 | 1026.6213
16753 | SOTS 7.5 | 11 | 11.56 | 1.09 | 2.8 | 22 | [5 empty] | 2018-12-04 17:00 | 24 | 12.34 | 12.1737 | 1.33 | 1.3121 | 5.1 | 5.0313 | 0.98737 | 1026.626
16754 | SOTS 7.5 | 12 | 12.34 | 1.33 | 5.1 | 24 | [5 empty] | 2018-12-13 17:00 | 26 | 11.04 | 10.913 | 1.125 | 1.1121 | 3.85 | 3.8057 | 0.98551 | 1026.5173
16755 | SOTS 7.5 | 13 | 11.06 | 0.92 | 2.6 | 26 | [5 empty] | 2018-12-22 17:00 | 28 | 7.65 | 7.6369 | 0.83 | 0.82858 | 1.3 | 1.2978 | 0.97599 | 1026.3546
16769 | SOTS 7.5 | 27 | 11.02 | 0.96 | 2.7 | 26 | -0.36298 |  |  |  |  | 2018-12-31 17:00 | 30 | 10.71 | 10.745 | 1.06 | 1.0635 | 2.5 | 2.5082 | 0.97103 | 1026.4738
16756 | SOTS 7.5 | 14 | 7.65 | 0.83 | 1.3 | 28 | [5 empty] | 2019-01-09 17:00 | 32 | 9.855 | 9.7045 | 0.92 | 0.90595 | 1.6 | 1.5756 | 0.98935 | 1026.4409
16757 | SOTS 7.5 | 15 | 10.71 | 1.06 | 2.5 | 30 | [5 empty] | 2019-01-18 17:00 | 34 | 9.2 | 9.1498 | 0.98 | 0.97465 | 1.3 | 1.2929 | 0.97969 | 1026.3383
16758 | SOTS 7.5 | 16 | 9.87 | 0.9 | 1.6 | 32 | [5 empty] | 2019-01-27 17:00 | 36 | 9.89 | 9.7298 | 0.9 | 0.88542 | 1.2 | 1.1806 | 0.99028 | 1026.4423
16768 | SOTS 7.5 | 26 | 9.84 | 0.94 | 1.6 | 32 | -0.30488 |  |  |  |  | 2019-02-05 17:00 | 38 | 9.74 | 9.5915 | 0.94 | 0.92567 | 1.7 | 1.6741 | 0.98936 | 1026.3983
16759 | SOTS 7.5 | 17 | 9.2 | 0.98 | 1.3 | 34 | [5 empty] | 2019-02-14 17:00 | 40 | 11.01 | 10.8498 | 1.15 | 1.1333 | 2.7 | 2.6607 | 0.98859 | 1026.4862
16760 | SOTS 7.5 | 18 | 9.89 | 0.9 | 1.2 | 36 | [5 empty] | 2019-02-23 17:00 | 42 | 9.13 | 9.0618 | 0.87 | 0.8635 | 1.3 | 1.2903 | 0.98156 | 1026.4504
16761 | SOTS 7.5 | 19 | 9.74 | 0.94 | 1.7 | 38 | [5 empty] | 2019-03-04 17:00 | 44 | 8.06 | 7.9339 | 0.8 | 0.78748 | 0.9 | 0.88591 | 0.9898 | 1026.3676
16762 | SOTS 7.5 | 20 | 11.01 | 1.15 | 2.7 | 40 | [5 empty] | 2019-03-13 17:00 | 46 | 10.41 | 10.2516 | 1.21 | 1.1916 | 3.2 | 3.1513 | 0.98922 | 1026.5092
16763 | SOTS 7.5 | 21 | 9.13 | 0.87 | 1.3 | 42 | [5 empty] | 2019-03-22 17:00 | 48 | 8.905 | 8.7645 | 1.095 | 1.0777 | 2.05 | 2.0177 | 0.98994 | 1026.3496
16764 | SOTS 7.5 | 22 | 8.06 | 0.8 | 0.9 | 44
16765 | SOTS 7.5 | 23 | 10.41 | 1.21 | 3.2 | 46
16766 | SOTS 7.5 | 24 | 8.91 | 1.09 | 2.1 | 48
16767 | SOTS 7.5 | 25 | 8.9 | 1.1 | 2 | 48 | -0.11236
CSIRO Hydrochemistry Laboratory
Date Issued: | 2019-08-28 |  |  | Castray Esplanade
Sample receipt date: | 09 AUG 2019 |  |  | Hobart TAS 7000
Data analysis date: | 14/08/2019 & SiO4 20/8/2019 |  |  | Email: O&AHydrochemistryFacility@csiro.au
Client Details: | Di Davies |  |  |  | Analysis Type:
| CSIRO |  |  |  | Nutrient |  | 28
Hydrochem SOP 001-4
Phone:
Email:
Notes: | NOx and Phosphate analysed on the 14 August. Silicate was also analysed on this date however there was large baseline drift at 660 nm. Silicate was reanalysed on the 20 august at 820 nm, where there was minimal baseline drift. Conclusion is that some turbidity or intereference at 660 nm wavelngth causing the baseline to drift upwards. The crm on the 14th August was Lot CD and was within 2% of the certified value for NOx and Phosphate. The crm on the 20th of August was Lot BY and was within 1 MDL of the certified value for silicate.
Christine Rees
Reference aterial for Nutrients in Seawater (RMNS)
| Standard range in µM was: |  |  |  |  | CRM Results | µM
[6 empty] | Lot CD NOx | 5.647 | 0.055 | certified value plus or minus expanded uncertainty
| Standard | NOx | PO4 | SiO4 |  | mean | 5.714
| Cal 0 | 0 | 0 | 0 |  | stdev | 0.008
| Cal 1 | 0.56 | 0.154 | 0.28 |  | 95% | 0.01568
| Cal 2 | 1.05 | 0.3 | 0.56 |  | Lot PO4 CD | 0.457 | 0.008 | certified value plus or minus expanded uncertainty
| Cal 3 | 2.1 | 0.6 | 1.12 |  | mean | 0.463
| Cal 4 | 4.2 | 1.2 | 2.1 |  | stdev | 0.002
| Cal 5 | 6.3 | 1.8 | 4.2 |  | 95% | 0.00392
| Cal 6 | 8.4 | 2.4 | 8.4 |  | Lot SiO4 BY | 1.805 | 0.064 | certified value plus or minus expanded uncertainty
| Cal 7 | 10.8 | 3 | 12.6 |  | mean | 1.625
| Cal 8 | 14 |  |  |  | stdev | 0.014
| Cal 9 | 21 |  |  |  | 95% | 0.02744
Method Detection Limits µM (MDL):
| NOx | 0.02
| PO4 | 0.02
| SiO4 | 0.2
CTD and Underway
[10 empty] | Latitude | Longitude | Time (UTC) | Nitrate (uM) |  | Phosphate (uM) | File |  | underway data ex Marlin
[9 empty] | underway | -46.6993 | 142.292 | 2018-10-05 20:00 | 12.6217 |  | 0.91636 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.6966 | 142.2928 | 2018-10-05 20:01 | 12.6048 |  | 0.91429 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.694 | 142.295 | 2018-10-05 20:03 | 12.623 |  | 0.9122 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.6918 | 142.298 | 2018-10-05 20:04 | 12.5022 |  | 0.90588 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.6894 | 142.3013 | 2018-10-05 20:06 | 12.5004 |  | 0.90975 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.6871 | 142.3046 | 2018-10-05 20:07 | 12.4174 |  | 0.90529 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.6847 | 142.3081 | 2018-10-05 20:09 | 12.3428 |  | 0.89856 | in2018_v04uwy016.SLK
[9 empty] | underway | -46.6823 | 142.3117 | 2018-10-05 20:10 | 12.3085 |  | 0.89622 | in2018_v04uwy016.SLK
IN2018_V04 Ellwood | [8 empty] | underway | -46.6799 | 142.3152 | 2018-10-05 20:12 | 12.2717 |  | 0.89657 | in2018_v04uwy016.SLK | [7 empty] | NOx | Phosphate | Silicate | date time
Deployment | RP | Time | Latitude | Longitude | Bottom depth | Pressure | Temperature (degC) | Temperature flag | Salinity (PSU) | Salinity flag | Oxygen (uM) | Oxygen flag | NOx (uM) | NOx flag | Phosphate (uM) | Phosphate flag | Silicate (uM) | Silicate flag | Ammonia (uM) | Ammonia flag | Nitrite (uM) | Nitrite flag | density kgm3 | umol Kg-1 | umol Kg-1 | umol Kg-1
24 | 36 | 2018-10-01T11:59:13Z | -47.0652 | 141.8745 | 4560 | 5.3 | 9.449 | 0 | 34.6528 | 0 | 282.883 | 0 | 12.68 | 0 | 0.96 | 0 | 2.5 | 0 | 0 | 63 | 0.247 | 0 | 1026.774 | 12.3494 | 0.93497 | 2.4348 | 2018-10-01 11:59
24 | 35 | 2018-10-01T11:59:10Z | -47.0652 | 141.8745 | 4560 | 6.5 | 9.444 | 0 | 34.6529 | 0 | 282.927 | 0 | 12.73 | 0 | 0.96 | 0 | 2.5 | 0 | -0.01 | 63 | 0.245 | 0 | 1026.775 | 12.398 | 0.93497 | 2.4348 | 2018-10-01 11:59
IN2018_V07 SOTS7.5 | [29 empty] | NOx | Phosphate | Silicate
SURVEY_NAME | STATION_NO | START_TIME | END_TIME | MIN_DEPTH | MAX_DEPTH | BOTTTOM_DEPTH | BOTTOM_TIME | BOTTOM_LAT | BOTTOM_LON | END_LAT | END_LON | START_LAT | START_LON | PROJECT_NAME | MARLIN_ID | MARLIN_UUID | BOTTLE_NUMBER | PRESSURE | NITRATE | NITRATE_QC | PHOSPHATE | PHOSPHATE_QC | SILICATE | SILICATE_QC | SALINITY | SALINITY_QC | TEMPERATURE | TEMPERATURE_QC | Density calc | umol Kg-1 | umol Kg-1 | umol Kg-1
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 25 | 5.2 |  | 141 |  | 141 |  | 141 |  | 141 | 10.026 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 26 | 5.8 | 11.605 | 0 | 0.877 | 0 | 2.412 | 0 |  | 141 | 10.026 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 27 | 5.9 | 11.598 | 0 | 0.885 | 0 | 2.442 | 0 | 34.764 | 0 | 10.026 | 0 | 1026.764 | 11.2957 | 0.86193 | 2.3783
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 24 | 29.4 | 11.598 | 0 | 0.888 | 0 | 2.459 | 0 | 34.766 | 0 | 10.037 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 23 | 29.8 | 11.618 | 0 | 0.886 | 0 | 2.466 | 0 |  | 141 | 10.037 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 22 | 29.9 |  | 141 |  | 141 |  | 141 |  | 141 | 10.038 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 19 | 50.4 |  | 141 |  | 141 |  | 141 |  | 141 | 10.036 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 20 | 50.5 | 11.602 | 0 | 0.887 | 0 | 2.481 | 0 |  | 141 | 10.036 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 21 | 51.3 | 11.586 | 0 | 0.886 | 0 | 2.491 | 0 | 34.765 | 0 | 10.036 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 18 | 100 | 11.577 | 0 | 0.886 | 0 | 2.454 | 0 | 34.766 | 0 | 10.05 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 17 | 126 | 11.573 | 0 | 0.883 | 0 | 2.469 | 0 | 34.768 | 0 | 10.059 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 16 | 150.9 | 11.577 | 0 | 0.88 | 0 | 2.448 | 0 | 34.768 | 0 | 10.063 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 15 | 175.9 | 11.582 | 0 | 0.891 | 0 | 2.455 | 0 | 34.768 | 0 | 10.064 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 14 | 200.7 | 11.562 | 0 | 0.881 | 0 | 2.468 | 0 | 34.768 | 0 | 10.066 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 13 | 225 | 11.597 | 0 | 0.887 | 0 | 2.473 | 0 | 34.767 | 0 | 10.063 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 12 | 250.7 | 11.638 | 0 | 0.891 | 0 | 2.496 | 0 | 34.766 | 0 | 10.058 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 11 | 301 | 15.343 | 0 | 1.091 | 0 | 4.134 | 0 | 34.722 | 0 | 9.6 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 10 | 351.2 | 15.683 | 0 | 1.116 | 0 | 4.247 | 0 | 34.688 | 0 | 9.384 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 9 | 481.3 | 16.353 | 0 | 1.17 | 0 | 4.489 | 0 | 34.615 | 0 | 8.923 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 8 | 600.3 | 19.696 | 0 | 1.373 | 0 | 7.029 | 0 | 34.531 | 0 | 8.291 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 7 | 800.7 | 26.634 | 0 | 1.807 | 0 | 17.533 | 0 | 34.431 | 0 | 6.7 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 6 | 999.7 | 30.48 | 0 | 2.086 | 0 | 31.078 | 0 | 34.37 | 0 | 5.032 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 5 | 1200.4 | 33.485 | 0 | 2.3 | 0 | 48.733 | 0 | 34.395 | 0 | 3.768 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 4 | 1402.5 | 34.851 | 0 | 2.413 | 0 | 67.585 | 0 | 34.493 | 0 | 3.255 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 3 | 1599.2 | 35.036 | 0 | 2.429 | 0 | 77.8 | 0 | 34.564 | 0 | 2.852 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 2 | 1800.7 | 34.414 | 0 | 2.386 | 0 | 81.554 | 0 | 34.626 | 0 | 2.575 | 0
IN2018_V07 | 1 | 2018-08-22 11:08 | 2018-08-22 13:07 | 5 | 2251 |  | 2018-08-22 11:54 |  |  |  |  | -47.0075 | 142.208 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | c10579e5-8867-406e-ac41-f8e679c7b1bc | 1 | 2000.1 | 33.742 | 0 | 2.345 | 0 | 85.475 | 0 | 34.674 | 0 | 2.414 | 0
IN2019_V02 SOTS8 | [29 empty] | NOx | Phosphate | Silicate
SURVEY_NAME | STATION_NO | START_TIME | END_TIME | MIN_DEPTH | MAX_DEPTH | BOTTTOM_DEPTH | BOTTOM_TIME | BOTTOM_LAT | BOTTOM_LON | END_LAT | END_LON | START_LAT | START_LON | PROJECT_NAME | MARLIN_ID | MARLIN_UUID | BOTTLE_NUMBER | PRESSURE | NITRATE | NITRATE_QC | PHOSPHATE | PHOSPHATE_QC | SILICATE | SILICATE_QC | SALINITY | SALINITY_QC | TEMPERATURE | TEMPERATURE_QC | Density calc | umol Kg-1 | umol Kg-1 | umol Kg-1
IN2019_V02 | 2 | 2019-03-16 13:31 | 2019-03-16 15:50 | 5.3 | 2001.5 | 4151.2 | 2019-03-16 14:28 | -46.9802 | 142.8131 | -46.9802 | 142.8131 | -46.9802 | 142.8131 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 23 | 5.3 | 8.963 | 0 | 0.732 | 0 | 0.845 | 0 | 34.576 | 0 | 11.04 | 0 | 1026.4386 | 8.7321 | 0.71315 | 0.82323
IN2019_V02 | 2 | 2019-03-16 13:31 | 2019-03-16 15:50 | 5.3 | 2001.5 | 4151.2 | 2019-03-16 14:28 | -46.9802 | 142.8131 | -46.9802 | 142.8131 | -46.9802 | 142.8131 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 5.5 | 8.957 | 0 | 0.746 | 0 | 0.833 | 0 | 34.576 | 0 | 11.04 | 0 | 1026.4386 | 8.7263 | 0.72678 | 0.81154
IN2019_V02 | 3 | 2019-03-17 01:02 | 2019-03-17 01:45 | 4.5 | 90.2 | 4618.5 | 2019-03-17 01:22 | -46.9912 | 142.6647 | -46.9912 | 142.6647 | -46.9912 | 142.6647 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 4.5 | 9.263 | 0 | 0.753 | 0 | 0.96 | 0 | 34.575 | 0 | 10.943 | 0 | 1026.4553 | 9.0243 | 0.73359 | 0.93526
IN2019_V02 | 7 | 2019-03-19 00:54 | 2019-03-19 01:35 | 7.3 | 99.3 | 4401.6 | 2019-03-19 01:14 | -46.8472 | 142.3457 | -46.8472 | 142.3457 | -46.8472 | 142.3457 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 7.3 | 9.326 | 0 | 0.742 | 0 | 1.021 | 0 | 34.57 | 0 | 10.859 | 0 | 1026.4665 | 9.0855 | 0.72287 | 0.99467
IN2019_V02 | 8 | 2019-03-19 11:59 | 2019-03-19 12:39 | 7 | 100.1 | 4448.3 | 2019-03-19 12:18 | -46.8278 | 142.3716 | -46.8278 | 142.3716 | -46.8278 | 142.3716 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 7.6 | 9.273 | 0 | 0.728 | 0 | 0.872 | 0 | 34.572 | 0 | 11.141 | 0 | 1026.4171 | 9.0343 | 0.70926 | 0.84956
IN2019_V02 | 12 | 2019-03-20 03:47 | 2019-03-20 06:00 | 11.1 | 2000.9 | 4515.7 | 2019-03-20 04:35 | -46.859 | 141.6433 | -46.859 | 141.6433 | -46.859 | 141.6433 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 11.1 | 8.792 | 0 | 0.723 | 0 | 0.775 | 0 | 34.574 | 0 | 11.463 | 0 | 1026.3596 | 8.5662 | 0.70443 | 0.7551
IN2019_V02 | 12 | 2019-03-20 03:47 | 2019-03-20 06:00 | 11.1 | 2000.9 | 4515.7 | 2019-03-20 04:35 | -46.859 | 141.6433 | -46.859 | 141.6433 | -46.859 | 141.6433 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 23 | 11.7 | 8.8 | 0 | 0.721 | 0 | 0.844 | 0 | 34.574 | 0 | 11.46 | 0 | 1026.3601 | 8.574 | 0.70248 | 0.82232
IN2019_V02 | 13 | 2019-03-24 01:22 | 2019-03-24 02:01 | 5.6 | 150.3 | 4085 | 2019-03-24 01:38 | -46.8526 | 141.8197 | -46.8526 | 141.8197 | -46.8526 | 141.8197 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 6.3 | 8.457 | 0 | 0.701 | 0 | 0.571 | 0 | 34.564 | 0 | 11.184 | 0 | 1026.4031 | 8.2395 | 0.68297 | 0.55631
IN2019_V02 | 15 | 2019-03-27 01:22 | 2019-03-27 01:56 | 10.5 | 120.4 | 4453.5 | 2019-03-27 01:36 | -46.96 | 142.1967 | -46.96 | 142.1967 | -46.96 | 142.1967 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 10.6 | 11.249 | 0 | 0.867 | 0 | 2.045 | 0 | 34.569 | 0 | 10.283 | 0 | 1026.5672 | 10.9579 | 0.84456 | 1.9921
IN2019_V02 | 18 | 2019-03-28 12:05 | 2019-03-28 12:41 | 9.4 | 100.5 | 4644 | 2019-03-28 12:20 | -46.7971 | 142.0293 | -46.7971 | 142.0293 | -46.7971 | 142.0293 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 10.3 | 13.341 | 133 | 0.988 | 133 | 3.023 | 133 | 34.601 | 133 | 10.464 | 0 | 1026.5606 | 12.9958 | 0.96244 | 2.9448
IN2019_V02 | 19 | 2019-03-29 00:58 | 2019-03-29 01:34 | 10.2 | 150.2 | 3769.6 | 2019-03-29 01:10 | -46.8631 | 142.1676 | -46.8631 | 142.1676 | -46.8631 | 142.1676 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 10.2 | 10.854 | 0 | 0.845 | 0 | 1.656 | 0 | 34.563 | 0 | 10.368 | 0 | 1026.5477 | 10.5733 | 0.82315 | 1.6132
IN2019_V02 | 21 | 2019-03-30 01:02 | 2019-03-30 02:07 | 9.6 | 301 | 4410.9 | 2019-03-30 01:15 | -46.8596 | 142.3048 | -46.8596 | 142.3048 | -46.8596 | 142.3048 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 10.7 | 10.822 | 0 | 0.844 | 0 | 1.507 | 0 | 34.563 | 0 | 10.315 | 0 | 1026.5569 | 10.542 | 0.82217 | 1.468
IN2019_V02 | 22 | 2019-03-30 12:00 | 2019-03-30 13:10 | 9.9 | 500.3 | 4336.2 | 2019-03-30 12:22 | -46.8549 | 142.3677 | -46.8549 | 142.3677 | -46.8549 | 142.3677 | SOTS: Southern Ocean Time Series automated moorings for climate and carbon cycle studies southwest of Tasmania |  | b9b978c4-410f-4d4c-9960-91582a997a95 | 24 | 10.8 | 11.176 | 0 | 0.857 | 0 | 1.766 | 0 | 34.573 | 0 | 10.301 | 0 | 1026.5672 | 10.8868 | 0.83482 | 1.7203

## Mclane comments
Hi Di,
Bottom Side Tubing: Polyurethane, Blue, Transparent 1/8" ID x 1/4" OD, 85D Shore A
We've actually had some recent trouble sourcing this. Lately we've been getting:
http://www.bimba.com/en/Products-and-Cad/Tubing/Inch/Polyurethane-Tubing/18-OD-X-116-ID-Polyurethane-Tubing/18-OD-x-116-ID-Polyurethane-Tubing2/?searchRef=yes&pn=PU-250-27
The bottom side tubing is less critical since it doesn't come in contact with the sample. The blue isn't necessary either, but transparent helps to identify bubbles.
Top Side Teflon Tubing: PFA Tubing Natural 1/8" OD x 1/16"
https://www.idex-hs.com/pfa-tubing-natural-1-8-od-x-1-16-id-x-50ft.html
We could of course supply some from here, but shipping will cost about as much or more. Let me know if you have any trouble getting something local.
Best regards,
Tim Shanahan
Just to be clear, you mean the cross sectional area is 1.51mm2 and hence the diameter is 1.39mm? Yes.

## inlet
Proved to be too difficult to obtain stock, inserted Cu:Ni tube at the end of the intake and also where the outlet exits the hose.
Note: need to use shoreA 84 or softer, designed not to be ridiculously difficult to insert. Rework station at 200C went well.
The corrosion rate at 200 days exposure in seawater (mil per year)
copper C10200 is about 1.6m/yr
Be:Cu C17200 is about 1.3m/yr
90:10 Cu Ni 1m/yr
70:30 Cu Ni with 0.5 Fe 0.4m/yr
Cu:Ni:Sn suggested, more has antifouling characteristics and low corrosion rates but no data yet.
Paul McMaster.
Managing Director
James Coppell Lee
20 Merri Concourse Campbellfield
Victoria. Australia. 3061.
Ph: +613 9357 9613
Fax: +613 9357 9608
Mbl: 0427 551 625
Email: sales@jacopplee.com
See the Prelaunch RAS intake.docx
dropbox and sofs7 files.

## phytoplankton
NO Glutaraldehyde, all HgCl2
RAS and mooring | Sample date time UTC | total wt g | net sample wt g. | sample analysis
RAS 9 SOFS7.5 | Event 1 of 48 at 08/27/18 16:00:00 | 428.86 | 418.553 | phyto
RAS 9 SOFS7.5 | Event 3 of 48 at 09/05/18 16:00:00 | 411.38 | 401.073 | phyto
RAS 9 SOFS7.5 | Event 5 of 48 at 09/14/18 16:00:00 | 430.54 | 420.233 | phyto
RAS 9 SOFS7.5 | Event 7 of 48 at 09/23/18 16:00:00 | 415.65 | 405.343 | phyto
RAS 9 SOFS7.5 | Event 9 of 48 at 10/02/18 16:00:00 | 423.34 | 413.033 | phyto
RAS 9 SOFS7.5 | Event 11 of 48 at 10/11/18 16:00:00 | 417.09 | 406.783 | phyto
RAS 9 SOFS7.5 | Event 13 of 48 at 10/20/18 16:00:00 | 424.97 | 414.663 | phyto
RAS 9 SOFS7.5 | Event 15 of 48 at 10/29/18 16:00:00 | 402.34 | 392.033 | phyto
RAS 9 SOFS7.5 | Event 17 of 48 at 11/07/18 16:00:00 | 370.48 | 360.173 | phyto
RAS 9 SOFS7.5 | Event 19 of 48 at 11/16/18 16:00:00 | 423.9 | 413.593 | phyto
RAS 9 SOFS7.5 | Event 21 of 48 at 11/25/18 16:00:00 | 379.69 | 369.383 | phyto
RAS 9 SOFS7.5 | Event 23 of 48 at 12/04/18 16:00:00 | 423.9 | 413.593 | phyto
RAS 9 SOFS7.5 | Event 25 of 48 at 12/13/18 16:00:00 | 350.13 | 339.823 | phyto
RAS 9 SOFS7.5 | Event 27 of 48 at 12/22/18 16:00:00 | 406.61 | 396.303 | phyto
RAS 9 SOFS7.5 | Event 29 of 48 at 12/31/18 16:00:00 | 407.08 | 396.773 | phyto
RAS 9 SOFS7.5 | Event 31 of 48 at 01/09/19 16:00:00 | 414.04 | 403.733 | phyto
RAS 9 SOFS7.5 | Event 33 of 48 at 01/18/19 16:00:00 | 421.9 | 411.593 | phyto
RAS 9 SOFS7.5 | Event 35 of 48 at 01/27/19 16:00:00 | 426.21 | 415.903 | phyto
RAS 9 SOFS7.5 | Event 37 of 48 at 02/05/19 16:00:00 | 421.86 | 411.553 | phyto
RAS 9 SOFS7.5 | Event 39 of 48 at 02/14/19 16:00:00 | 414.13 | 403.823 | phyto
RAS 9 SOFS7.5 | Event 41 of 48 at 02/23/19 16:00:00 | 417.93 | 407.623 | phyto
RAS 9 SOFS7.5 | Event 43 of 48 at 03/04/19 16:00:00 | 396.02 | 385.713 | phyto
RAS 9 SOFS7.5 | Event 45 of 48 at 03/13/19 16:00:00 | 408.25 | 397.943 | phyto
RAS 9 SOFS7.5 | Event 47 of 48 at 03/22/19 16:00:00 | 399.77 | 389.463 | phyto
Decant 50mL for coco 0.8um nuclepore, mq rinse.
Preconcentrate all of the remaining volume with alkaline Lugols.

## build comments
RAS Orpheus loses data when the batteries are removed. There is no backup battery in this particular model.
Frame and plate hole alignments marginal and some needed redrilling, particularly four bolts in the top of the frame, possibly caused by warping during fabriaction of the frame
Holes in platic need to be well oversized
The radius in the corners of the plastic plates need to be bigger, sitting on welds otherwise.
Cut out for the valve is too small, difficult access especially home ports. This was done for strength with no consideration for access.
Furtherest jar (49:) should be possible to move it in, redraw it.
Water jet has not cut holes cleanly, all holes need to be checked before progressing, jars didn’t fit.
Jar spacing is irregular and too tight to run the bottom tubes between them
Jar heights are irregular and need cushioning underneath, otherwise the tie downs don’t fit and jars can rotate. Fitted better after packing up the tie down stand offs.
Tie down spacers/stand offs are too short. Washers to take up the differences before it goes out again. Done
The pump needs to be between the valve and the point above the load cell for access and reduce the run of tube to the chain. Repositioned. Not as shown in the drawing
The build ended up heavy at 180kg. Stiffness and weight reduction can be achived by fabricating out of angle?
The controller could not be removed after all stiffners were in place making access to the tubes beside it difficult. If it had a spacer in its bracket that could be knocked out, the controller could be removed during prep. The wider end fitting does not allow for sliding the controller out of the bracket. The Bolts coming through the bracket from the top are too long and cannot be removed.
A second set of tie down bars to stack ontop of the first set would do a better job of keeping the tubing flat. Easier than cable tying maybe.
Little choice in material for elbows, prefer not to use Kynar again, a very long lead time for anything else
Pattern of tie down bar holes needs to be alternate spacing, otherwise tilts when fastened. Additional tie down stand offs installed, not as show in drawing.
Mclane jars were super tight and need a purpose built lid opener. Workshop fabriacated a pair of spanners on second attempt but the material was not stiff enough, improved by gluing together. Serviceable
Tedlar bags had faults (increasing problem) and full of fibres. All bags were washed. Washer attaches to the MQ outlet with a Tee and syringe for emptying.
I have yet to work out why I had so much trouble filling the poison reservoir
Threads on the corner posts of the frame were not cut deep enough.
Now has lifting points - post installed 8mm 316 ss eyebolts (PJ's)
bottom of acrylic tubes failing, about 25%, new tubes. Contact McLane, order more. Try solvent and uncured UW epoxy?
The version depolyed was drawing version 7 that tries to address problems from SOFS7.
http://citeseerx.ist.psu.edu/viewdoc/download?doi=10.1.1.629.8265&rep=rep1&type=pdf
Note the bottom od is 62mm and the tube id 62.2 mm milled slightly off centre so that the glue line is thin and often dry because of the offset.
Workshop provided:
1 | DCM glue with original clearences
2 | extra turned out leaving a shoulder to keep the bottom centre with a glue line of 0.30mm
3 | extra turned out leaving a shoulder to keep the bottom centre with a glue line of 0.35 mm

## hardware
Left over from Orpheus IN2018_V02 build - lifting bridle stored in base.

## CO2
file: Carbon Results RAS9 SOFS7.5 2019.xlsx | [13 empty] | SOMMA QC data, RAS 2019 | [18 empty] | TA QC data, RAS 2019
Sample | Date_Time | RP/S# | Alkalinity | Alkalinity | SOMMA | TCO2 | TCO2 |  | Alkalinity | TCO2 |  |  |  | type | Bottle | Stn | Cast | Nisk | Depth | S temp | salt | counts | RT | TCO2 | Calfact | blank | TCT | last CRM | cert CRM | CRM sal | cert sal | batch | CRMs
type | yyyymmdd_hhmm | no | µmol/kg | Flag | salinity | µmol/kg | Flag | diln | µmol/kg | µmol/kg |  |  |  | CRM | 33 |  |  |  |  | 22.05 | 33.378 | 201418 | 8 | 2016.83 | 1.0029 | 5 | 7.2 | 2016.83 | 2017.95 | 33.378 | 33.357 | 153 | Alkalinity | µmol/kg
RAS9 SOFS7.5 | 2018-08-27 17:00 | 2 | 2257.768 | 2 | 34.142 | 2348.395 | 2 | 0.9835 | 2295.64 | 2387.7871 |  |  |  | CRM | 033a |  |  |  |  | 22.03 | 33.356 | 201363 | 9 | 2016.22 | 1.0029 | 5 | 21.1 | 2016.22 | 2017.95 | 33.356 | 33.357 | 153 | average | 2192.0286
RAS9 SOFS7.5 | 2018-09-05 17:00 | 4 | 2267.1525 | 2 | 34.264 | 2416.535 | 2 | 0.98892 | 2292.5607 | 2443.6173 | [22 empty] | stdev | 1.0118
RAS9 SOFS7.5 | 2018-09-14 17:00 | 6 | 2269.6756 | 2 | 34.323 | 2323.855 | 2 | 0.99134 | 2289.4939 | 2344.1463 | [22 empty] | n | 7
RAS9 SOFS7.5 | 2018-09-23 17:00 | 8 | 2265.0444 | 2 | 34.323 | 2404.475 | 2 | 0.98953 | 2289.0127 | 2429.9187 | [22 empty] | UCL | 2195.064
RAS9 SOFS7.5 | 2018-10-02 17:00 | 10 | 2263.2193 | 2 | 34.215 | 2365.575 | 2 | 0.98826 | 2290.1146 | 2393.6867 | [9 empty] | average | 33.367 |  | average | 2016.525 | [8 empty] | UWL | 2194.0522
RAS9 SOFS7.5 | 2018-10-11 17:00 | 12 | 2177.9366 | 2 | 33.738 | 2244.645 | 2 | 0.97807 | 2226.772 | 2294.9762 | [9 empty] | s.d. | 0.015556 |  | s.d. | 0.43134 | [8 empty] | LWL | 2190.0049
RAS9 SOFS7.5 | 2018-10-20 17:00 | 14 | 2265.53 | 2 | 34.202 | 2123.975 | 2 | 0.98978 | 2288.9259 | 2145.909 | [9 empty] | n | 2 |  | n | 2 | [8 empty] | LCL | 2188.9931
RAS9 SOFS7.5 | 2018-10-29 17:00 | 16 | 2258.6149 | 2 | 34.146 | 2347.265 | 2 | 0.98846 | 2284.9939 | 2374.6793 |  |  |  | SOMMA CRMs | µmol/kg | [9 empty] | CRMs | Average | cert CRM | Correction | Diff last-first |  |  |  | certified value | 2207.64
RAS9 SOFS7.5 | 2018-11-07 17:00 | 18 | 2215.7381 | 2 | 34.191 | 2240.965 | 2 | 0.98932 | 2239.651 | 2265.1502 |  |  |  | Control limit table | [9 empty] | 33 | 2016.83 | [7 empty] | Correction applied to TA data | 1.0071
RAS9 SOFS7.5 | 2018-11-16 17:00 | 20 | 2217.5382 | 2 | 34.197 | 2323.975 | 2 | 0.98924 | 2241.648 | 2349.2419 |  |  |  | average | 2016.525 | [8 empty] | 033a | 2016.22 | 2016.525 | 2017.95 | 1.425 | -0.61
RAS9 SOFS7.5 | 2018-11-25 17:00 | 22 | 2202.3864 | 2 | 34.171 | 2337.845 | 2 | 0.98902 | 2226.8459 | 2363.8089 |  |  |  | stdev | 0.43134
RAS9 SOFS7.5 | 2018-12-04 17:00 | 24 | 2254.667 | 2 | 34.105 | 2188.985 | 2 | 0.98737 | 2283.5106 | 2216.9883 |  |  |  | n | 2
RAS9 SOFS7.5 | 2018-12-13 17:00 | 26 | 2206.5072 | 2 | 34.026 | 2276.575 | 2 | 0.98551 | 2238.957 | 2310.0552 |  |  |  | UCL | 2017.819
RAS9 SOFS7.5 | 2018-12-22 17:00 | 28 | 2176.1221 | 2 | 33.742 | 2194.425 | 2 | 0.97599 | 2229.6512 | 2248.4044 |  |  |  | UWL | 2017.3877
RAS9 SOFS7.5 | 2018-12-31 17:00 | 30 | 2182.1546 | 2 | 33.544 | 2228.955 | 2 | 0.97103 | 2247.2472 | 2295.4436 |  |  |  | LWL | 2015.6623
RAS9 SOFS7.5 | 2019-01-09 17:00 | 32 | 2252.0996 | 2 | 34.191 | 2315.975 | 2 | 0.98935 | 2276.3523 | 2340.9155 |  |  |  | LCL | 2015.231
RAS9 SOFS7.5 | 2019-01-18 17:00 | 34 | 2241.5015 | 2 | 33.908 | 2163.945 | 2 | 0.97969 | 2287.9736 | 2208.8091 |  |  |  | Certified value | 2017.95
RAS9 SOFS7.5 | 2019-01-27 17:00 | 36 | 2225.5622 | 2 | 34.204 | 2153.145 | 2 | 0.99028 | 2247.4053 | 2174.2773
RAS9 SOFS7.5 | 2019-02-05 17:00 | 38 | 2269.4862 | 2 | 34.225 | 2214.635 | 2 | 0.98936 | 2293.8819 | 2238.4411
RAS9 SOFS7.5 | 2019-02-14 17:00 | 40 | 2258.4026 | 2 | 34.122 | 2233.785 | 2 | 0.98859 | 2284.4799 | 2259.5781
RAS9 SOFS7.5 | 2019-02-23 17:00 | 42 | 2193.1618 | 2 | 33.892 | 2176.635 | 2 | 0.98156 | 2234.3563 | 2217.5192 |  |  |  | SOMMA salinities | µmol/kg
RAS9 SOFS7.5 | 2019-03-04 17:00 | 44 | 2232.2406 | 2 | 34.217 | 2098.445 | 2 | 0.9898 | 2255.2434 | 2120.0691 |  |  |  | Control limit table
RAS9 SOFS7.5 | 2019-03-13 17:00 | 46 | 2263.129 | 2 | 34.179 | 2089.345 | 2 | 0.98922 | 2287.7805 | 2112.1035 |  |  |  | average | 33.367
RAS9 SOFS7.5 | 2019-03-22 17:00 | 48 | 2264.7505 | 2 | 34.217 | 2115.125 | 2 | 0.98994 | 2287.764 | 2136.6181 |  |  |  | stdev | 0.015556
[14 empty] | n | 2
file: CO2 results IN2018_V07_IN2018_V04.xlsx | [13 empty] | UCL | 33.4137
Sample | Date_Time | CTD | RP/S# | Depth | CTD | T_insitu | Latitude | Longitude | Alkalinity | Alkalinity | SOMMA | TCO2 | TCO2 | UWL | 33.3981
type | yyyymmdd_hhmm | no | no | db | Salinity | °C | °S | °E | µmol/kg | Flag | salinity | µmol/kg | Flag | LWL | 33.3359
IN2018_V07 | 2018-08-22 10:52 | 2 | 1 | 2000.1 | 34.675 | 2.414 | -47.0087 | 142.2124 | 2353.782 | 2 | 34.674 | 2260.4345 | 2 | LCL | 33.3203
IN2018_V07 | 2018-08-22 10:52 | 2 | 2 | 1599.2 | 34.567 | 2.852 | -47.0087 | 142.2124 | 2346.0874 | 2 | 34.566 | 2258.5537 | 2 | Certified value | 33.357
IN2018_V07 | 2018-08-22 10:52 | 2 | 3 | 1200.4 | 34.391 | 3.768 | -47.0087 | 142.2124 | 2315.9943 | 6 | 34.419 | 2221.739 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 4 | 800.7 | 34.433 | 6.7 | -47.0087 | 142.2124 | 2291.8209 | 2 | 34.434 | 2173.5998 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 5 | 481.3 | 34.619 | 8.923 | -47.0087 | 142.2124 | 2293.5349 | 2 | 34.601 | 2132.8335 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 6 | 301 | 34.725 | 9.6 | -47.0087 | 142.2124 | 2297.0258 | 2 | 34.734 | 2130.9927 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 7 | 225 | 34.711 | 10.063 | -47.0087 | 142.2124 | 2301.3095 | 2 | 34.776 | 2114.8263 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 8 | 175.9 | 34.772 | 10.064 | -47.0087 | 142.2124 | 2298.847 | 2 | 34.795 | 2115.1364 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 9 | 126 | 34.772 | 10.059 | -47.0087 | 142.2124 | 2299.9864 | 2 | 34.806 | 2115.2464 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 10 | 51.3 | 34.769 | 10.036 | -47.0087 | 142.2124 | 2301.2209 | 2 | 34.777 | 2113.8459 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 11 | 29.4 | 34.769 | 10.037 | -47.0087 | 142.2124 | 2301.6326 | 2 | 34.727 | 2115.0464 | 2
IN2018_V07 | 2018-08-22 10:52 | 2 | 12 | 5.9 | 34.768 | 10.026 | -47.0087 | 142.2124 | 2302.1243 | 2 | 34.78 | 2114.8363 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 6 | 2000 | 34.678 | 2.415 | -47.055 | 141.871 | 2352.8995 | 2 | 34.684 | 2256.242 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 9 | 1600 | 34.564 | 2.786 | -47.055 | 141.871 | 2341.1113 | 2 | 34.596 | 2254.2212 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 11 | 1200 | 34.409 | 3.821 | -47.055 | 141.871 | 2318.7186 | 2 | 34.439 | 2223.2688 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 14 | 800 | 34.431 | 6.779 | -47.055 | 141.871 | 2290.8805 | 2 | 34.435 | 2171.198 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 17 | 480 | 34.602 | 8.814 | -47.055 | 141.871 | 2292.6542 | 2 | 34.626 | 2132.5025 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 20 | 300 | 34.675 | 9.339 | -47.055 | 141.871 | 2294.5194 | 2 | 34.701 | 2134.7434 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 23 | 225 | 34.665 | 9.53 | -47.055 | 141.871 | 2296.6548 | 2 | 34.669 | 2114.4653 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 26 | 175 | 34.662 | 9.512 | -47.055 | 141.871 | 2295.1842 | 2 | 34.695 | 2113.9751 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 28 | 125 | 34.666 | 9.475 | -47.055 | 141.871 | 2296.1959 | 2 | 34.667 | 2115.3656 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 31 | 50 | 34.657 | 9.477 | -47.055 | 141.871 | 2296.8732 | 2 | 34.677 | 2114.4553 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 33 | 30 | 34.665 | 9.459 | -47.055 | 141.871 | 2296.0066 | 2 | 34.627 | 2117.7266 | 2
IN2018_V04 | 2018-10-01 07:51 | 24 | 35 | 5 | 34.655 | 9.444 | -47.055 | 141.871 | 2295.925 | 2 | 34.65 | 2114.7054 | 2
IN2019_V02
IN2019_V02 | 8 | 24 | 7.6 | 2019-03-19T11:58:59Z | 34.576 | 11.141 | 46.8283 | 142.3716 | 2294.4468 | 2 |  | 2092.4412 | 2
IN2019_V02 | 8 | 20 | 20.7 | 2019-03-19T11:58:59Z | 34.575 | 10.986 | 46.8283 | 142.3716 | 2295.893 | 2 |  | 2091.2207 | 2
IN2019_V02 | 13 | 24 | 6.3 | 2019-03-24T01:22:14Z | 34.567 | 11.184 | 46.8526 | 141.8199 | 2295.1525 | 2 |  | 2085.618 | 2
IN2019_V02 | 13 | 20 | 20.1 | 2019-03-24T01:22:14Z | 34.567 | 11.184 | 46.8526 | 141.8199 | 2295.9698 | 2 |  | 2088.1742 | 2
Note on K45: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K46: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K47: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K48: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K49: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K50: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K51: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K52: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K53: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K54: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K55: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis
Note on K56: Berry, Kate (O&A, Hobart): Berry, Kate (O&A, Hobart): Note: SOMMA salinity was used for TA analysis

## netcdf
deploymentyearstart | site | depth_nominal | remoteaccesssampler | metadata |  | sample | sample_qc | pressurerel | pressurerel_qc | temperature | temperature_qc | salinity | salinity_qc | time | weight | weight_qc | NOxconcentration | NOx_qc | phosphateconcentration | phosphate_qc | silicateconcentration | silicate_qc | totalalkalinity | totalalkalinity_qc | totalcarbondioxide | totalcarbondioxide_qc
deployment | SOTS7.5 |  |  | standard_name |  |  |  | sea_water_pressure_due_to_sea_water |  | sea_water_temperature |  | sea_water_practical_salinity |  |  |  |  | moles_of_nitrate_and_nitrite_per_unit_mass_in_sea_water |  | moles_of_phosphate_per_unit_mass_in_sea_water |  | moles_of_silicate_per_unit_mass_in_sea_water |  | moles_of_alkalinity_per_unit_mass_in_sea_water |  | moles_of_inorganic_carbon_per_unit_mass_in_sea_water
IN2018_V02 |  | depth nominal |  | long_name |  | sample number |  | sea_water_pressure_due_to_sea_water |  | sea_water_temperature |  | sea_water_practical_salinity |  | time of sample start | sample mass |  | moles_of_nitrate_and_nitrite_per_unit_mass_in_sea_water |  | moles_of_phosphate_per_unit_mass_in_sea_water |  | moles_of_silicate_per_unit_mass_in_sea_water |  | moles_of_alkalinity_per_unit_mass_in_sea_water |  | moles_of_inorganic_carbon_per_unit_mass_in_sea_water
SOFS-7.5-2018 |  | m |  | units |  | sampler position |  | dbar |  | oC |  | 1 |  | yyyy:mm:dd hh:mm:ss | kg |  | umol.kg-1 |  | umol.kg-1 |  | umol.kg-1 |  | umol.kg-1 |  | umol.kg-1
|  | nominal |  | uncertainty |  |  |  | 3 |  | 0.005 |  | 0.00054 |  |  | 1e-05 |  | 0.055 |  | 0.026 |  | 0.2 |  | 1.01 |  | 0.43
|  |  |  | comment | sample pairs, nutrients and separate sample for phytoplankton reported elsewhere (SOTS Phytoplankton abundance and biovolume). | [8 empty] | UTC
comment
|  |  |  | comment_method | TCO2 sample preservation issue flag 3 |  |  |  |  | Sea-Bird Electronics SBE37SM 7408 |  | Sea-Bird Electronics SBE37SM 7408 |  |  |  |  | FSA |  | FSA |  | FSA |  | potentiometric |  | coulometric
comment_sample
|  |  |  | comment_QC_report | Davies, Diana M., Jansen, Peter, Trull, Thomas W. IMOS - ABOS Southern Ocean Time Series (SOTS) - Quality Assessment and Control Report - Remote Access Sampler: Sample Analysis http://dx.doi.org/10.26198/5e156a63a8f75
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 2 | 1 | 4.5 | 1 | 9.793 | 1 | 34.7147 | 1 | 2018-08-27 17:00 | 0.41493 | 1 | 12.0169 | 1 | 0.95561 | 3 | 3.0698 | 1 | 2295.64 | 1 | 2387.7871 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 4 | 1 | 4.5 | 1 | 9.4969 | 1 | 34.648 | 1 | 2018-09-05 17:00 | 0.33982 | 1 | 12.6257 | 1 | 1.0242 | 3 | 2.5606 | 1 | 2292.5607 | 1 | 2443.6173 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 6 | 1 | 4.5 | 1 | 9.4426 | 1 | 34.6227 | 1 | 2018-09-14 17:00 | 0.38417 | 1 | 12.5163 | 1 | 1.0905 | 3 | 2.5544 | 1 | 2289.4939 | 1 | 2344.1463 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 8 | 1 | 4.5 | 1 | 9.6511 | 1 | 34.6862 | 1 | 2018-09-23 17:00 | 0.3963 | 1 | 12.2734 | 1 | 0.98423 | 3 | 2.3622 | 1 | 2289.0127 | 1 | 2429.9187 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 10 | 1 | 4.5 | 1 | 9.3273 | 1 | 34.6216 | 1 | 2018-10-02 17:00 | 0.41032 | 1 | 12.7819 | 1 | 1.0052 | 3 | 2.2666 | 1 | 2290.1146 | 1 | 2393.6867 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 12 | 1 | 4.5 | 1 | 8.9429 | 1 | 34.4945 | 1 | 2018-10-11 17:00 | 0.39677 | 1 | 13.3437 | 1 | 1.0655 | 3 | 3.2861 | 1 | 2226.772 | 1 | 2294.9762 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 14 | 1 | 4.5 | 1 | 9.6268 | 1 | 34.5552 | 1 | 2018-10-20 17:00 | 0.41262 | 1 | 12.5765 | 1 | 1.2104 | 3 | 3.7395 | 1 | 2288.9259 | 1 | 2145.909 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 16 | 1 | 4.5 | 1 | 9.1446 | 1 | 34.5448 | 1 | 2018-10-29 17:00 | 0.40373 | 1 | 13.6074 | 1 | 1.212 | 3 | 3.7442 | 1 | 2284.9939 | 1 | 2374.6793 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 18 | 1 | 4.5 | 1 | 9.1436 | 1 | 34.56 | 1 | 2018-11-07 17:00 | 0.41862 | 1 | 11.6461 | 1 | 0.93523 | 3 | 3.0518 | 1 | 2239.651 | 1 | 2265.1502 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 20 | 1 | 4.5 | 1 | 9.4662 | 1 | 34.5688 | 1 | 2018-11-16 17:00 | 0.41159 | 1 | 10.7811 | 1 | 0.94519 | 3 | 2.7568 | 1 | 2241.648 | 1 | 2349.2419 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 22 | 1 | 4.5 | 1 | 9.9175 | 1 | 34.5505 | 1 | 2018-11-25 17:00 | 0.40317 | 1 | 11.3853 | 1 | 1.0735 | 3 | 2.7577 | 1 | 2226.8459 | 1 | 2363.8089 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 24 | 1 | 4.5 | 1 | 9.848 | 1 | 34.5413 | 1 | 2018-12-04 17:00 | 0.4159 | 1 | 12.1737 | 1 | 1.3121 | 3 | 5.0313 | 3 | 2283.5106 | 1 | 2216.9883 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 26 | 1 | 4.5 | 1 | 10.4151 | 1 | 34.5264 | 1 | 2018-12-13 17:00 | 0.4219 | 1 | 10.913 | 1 | 1.1121 | 3 | 3.8057 | 1 | 2238.957 | 1 | 2310.0552 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 28 | 1 | 4.5 | 1 | 11.5288 | 1 | 34.572 | 1 | 2018-12-22 17:00 | 0.41155 | 1 | 7.6369 | 1 | 0.82858 | 3 | 1.2978 | 1 | 2229.6512 | 1 | 2248.4044 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 30 | 1 | 4.5 | 1 | 10.7401 | 1 | 34.5446 | 1 | 2018-12-31 17:00 | 0.41215 | 1 | 10.745 | 1 | 1.0635 | 3 | 2.5082 | 1 | 2247.2472 | 1 | 2295.4436 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 32 | 1 | 4.5 | 1 | 10.9866 | 1 | 34.5592 | 1 | 2019-01-09 17:00 | 0.40382 | 1 | 9.7045 | 1 | 0.90595 | 3 | 1.5756 | 1 | 2276.3523 | 1 | 2340.9155 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 34 | 1 | 4.5 | 1 | 11.762 | 1 | 34.611 | 1 | 2019-01-18 17:00 | 0.32724 | 1 | 9.1498 | 1 | 0.97465 | 3 | 1.2929 | 2 | 2287.9736 | 1 | 2208.8091 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 36 | 1 | 4.5 | 1 | 10.8936 | 1 | 34.5397 | 1 | 2019-01-27 17:00 | 0.40762 | 1 | 9.7298 | 1 | 0.88542 | 3 | 1.1806 | 2 | 2247.4053 | 1 | 2174.2773 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 38 | 1 | 4.5 | 1 | 11.3633 | 1 | 34.5929 | 1 | 2019-02-05 17:00 | 0.41426 | 1 | 9.5915 | 1 | 0.92567 | 3 | 1.6741 | 2 | 2293.8819 | 1 | 2238.4411 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 40 | 1 | 4.5 | 1 | 10.5432 | 1 | 34.516 | 1 | 2019-02-14 17:00 | 0.38571 | 1 | 10.8498 | 1 | 1.1333 | 3 | 2.6607 | 1 | 2284.4799 | 1 | 2259.5781 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 42 | 1 | 4.5 | 1 | 10.8014 | 1 | 34.5286 | 1 | 2019-02-23 17:00 | 0.37304 | 1 | 9.0618 | 1 | 0.8635 | 3 | 1.2903 | 2 | 2234.3563 | 1 | 2217.5192 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 44 | 1 | 4.5 | 1 | 11.4328 | 1 | 34.5696 | 1 | 2019-03-04 17:00 | 0.39794 | 1 | 7.9339 | 1 | 0.78748 | 3 | 0.88591 | 2 | 2255.2434 | 1 | 2120.0691 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 46 | 1 | 4.5 | 1 | 10.5682 | 1 | 34.5513 | 1 | 2019-03-13 17:00 | 0.40912 | 1 | 10.2516 | 1 | 1.1916 | 3 | 3.1513 | 1 | 2287.7805 | 1 | 2112.1035 | 3
2018 |  | 4.5 | RAS3-48-500FH serial no: 14384-01 Orpheus |  |  | 48 | 1 | 4.5 | 1 | 11.5099 | 1 | 34.5647 | 1 | 2019-03-22 17:00 | 0.38946 | 1 | 8.7645 | 1 | 1.0777 | 3 | 2.0177 | 1 | 2287.764 | 1 | 2136.6181 | 3